Design of Neutron Instrumentation
via Simulation
E. Farhi,
Institut Laue-Langevin, Grenoble, France.
<www.mcstas.org> project supported by EU FP4/5/6/7 since 1996

---
Introducing the Neutron
Electrically Neutral
penetrate deeply into matter – non destructive ;
can not be handled with E/B fields.
Microscopically Magnetic (polarized neutron beams)
they possess a magnetic dipole moment ;
spin is sensitive to B fields, thus can do NMR-like studies.
Ångstrom wavelengths
neutron wavelengths range from 0.1 Å to 1000 Å ;
well suited for studying atomic/molecular structure.
Energies of milli-electronvolts (1 µeV - 1eV)
same magnitude as most motions in solids and liquids ;
suited to studying phonons and magnons.
Randomly sensitive (cross sections)
well suited to distinguish light atoms (CNOH) and isotopes
within materials.
A fancy neutron

---
Usage of neutrons: instruments for users
About 5000 users in Europe
in Physics, Chemistry, Biology, and Nuclear Physics
14 neutron facilities – 250 instruments in Europe
within which ILL (started 1971) is the world's most intense.
Neutrons generated in nuclear reactors (ILL, Munich, LLB)
and spallation sources (accelerator based, ISIS, PSI)
Complementary with X-rays
not sensitive to the same features, but same wavelength range
Used to measure structure and dynamics of materials
diffractometers (powders, liquids, crystals) - structure
inelastic scattering spectrometers (powders, liquids, crystals) - dynamics
small angle machines for large scale structures (polymers, colloids, gels)
magnetic materials structure and dynamics using polarized neutrons (spin)
nuclear and fundamental physics with neutrons (fission, neutron properties, g)

---
Basics of neutron scattering
Neutrons interact with matter ...
Material
2q
Ki, Ei
Q, w
(d=2p/Q)
Kf, Ei
Kf=Ki+Q
Ef=Ei+w
Selections rules
Bragg's law (diffraction on structure – atoms separated by distance d)
%lambda = 2 %pi/K_i = 2 d sin %theta
Scattering law (intensity per solid angle and energy, dynamics)
{d^2 %sigma} over {d%OMEGA d%E_f} = %sigma over {4%pi} K_f over K_i S(Q,%omega)
Dynamical structure factor S(Q,w) is characteristic for each material
Reflects ordering of matter (atom/molecule positions – movements - domains)

---
Interaction processes
Elastic scattering: no energy exchange with matter
w = 0, |Ki|=|Kf|
Ruled by the Bragg's law → structure
Inelastic scattering: energy exchange with matter
Neutron can gain or loose energy with is transfered to matter
Couples e.g. to matter dynamics (phonons, magnons, …)
Coherent scattering: atoms/molecules in matter vibrates coherently
e.g. Structure (elastic), sound waves, spin waves, ...
Incoherent scattering: atoms/molecules have random processes in matter.
e.g. Isotopes, defects, heat diffusion, ...
S(Q) = int S(Q,%omega) d%omega

---
Example of materials: powder and liquid
Structure
S(Q)
Elastic scattering
(diffractometer)
Structure peaks
(atom/molecule spacing)
Dynamics
S(Q,w)
Phonons
(atom/molecule motions)
Inelastic scattering
(time-of-flight)

---
Why simulate neutron instruments ?
Help instrument responsibles
instrument optimization and design
Help scientists
plan experiments (for proposal application: is the experiment feasible ?)
guide experiments (during beam time allocation: optimize beam usage)
help understand results and analyze data (after experiment: what is this ?)
How to proceed ?
Install a Monte Carlo simulation software for neutron instrumentation
simulation: McStas, VitESS, NISP, ResTrax
Collect information about instrument geometry and configuration
Collect information about the sample/measurement to be performed
Use existing instrument model/template and customize it
or write model from scratch
Launch simulation
Compare with what you (think) know/measure

---
Presentation of McStas
Flexible, general simulation utility for neutron scattering experiments.
Original design for Monte Carlo Simulation of triple axis spectrometers.
Developed at RISØ and ILL (started 1998).
Mainly funded by EU (FP 4-7): 2.5+1 people full time, plus projects.
Open source, GPL2. Works on all systems, binaries for Linux, Win$, MacOSX
Most widely used code in its ecological niche (compared to other similar codes)..
Probably a few hundred users worldwide (160 registered)
some contributors (community based).
Has been used to model most existing and future
neutron scattering instruments.
mcstas-users@mcstas.org mailing list
mcstas-support@mcstas.org developer contact
<www.mcstas.org>

---
Main capabilities of McStas
McStas requires an instrument description (text)
Many instrument examples are given with
the software, covering all instrument
classes (TOF, TAS, SANS, DIF, ...).
McStas creates an instrument executable
model for which configuration parameters can be changed at will.
Exe
Exe
Parameters
Parameters
Results
Simulation
results
Plots
Detector Plots
3D view
3D view

---
More about McStas
McStas is a compiler. It translates a clean instrument description
(using a dedicated grammar) into a flat C file, which is compiled
and executed.
Components used instruments are pieces of C code, with dedicated
grammar and function calls to ease programming.
Creating new instruments and components is damn easy.
Documentation:
Extensive code documentation
User and Component manuals,
mailing list/forum, tutorial,
50 instrument examples.
McStas has built-in support for cluster
and grid computing.
Simulation = Source+guide+primary+sample+secondary

---
Source: Generating neutron events
Neutron events may be currently generated from:
A black body (Maxwellian distributions)
Measured flux/divergence distributions
MCNP and Tripoli (nuclear reactor simulation codes) event files
Other neutron generators are in principle easy to write
(e.g. from GEANT, FLUKA, ...)
top
 n = (R, v, t, s, p)
side

---
Guide systems: bring neutrons to instruments
Guide systems are usually curved
to remove g rays from the core.
Here is a 3D view generated
from the H15 guide, feeding
the D11 instrument (polymers).
Sections: 3x20 cm, length 100 m
Optimize:
Guide shape (curvature, section)
Guide reflectivity (coating):
Ni, Ti
R=1 for a ~ 0.1 l

---
Neutron optics: monochromators (primary spect.)
Monochromators are used to extract a
sharp neutron energy distribution from a white beam.
Rely on Bragg's law.
Use single crystal assembly, with focusing geometry.
Size: Typically 20x20 cm
Design and characterisation are eased using simulation,
intensity vs. curvature radius,
thickness of blades
rocking curve (reflectivity)
scattering from the mechanics (generated background)
choice of crystals
Cu, Ge, Si, C (graphite)
CuMnAl, CoFe (pol.)

---
Neutron optics: choppers (primary spect.)
3 beams (monk.)
200 Al blades
Gd absorber
Replaced with NiTi
Rotation 10000 rpm
Here, we change the Fermi chopper coating to increase the flux at the sample position.
Disk choppers
Coating
Angle

---
Sample environments: hot/cold, magnets
Sample is usually studied at low or high temperatures. Cryostats and furnaces
Must not generate too much background (all materials in the beam scatter, not only sample)..
TOF E0=33 meV. Liquid Rb.
Coherent single
scattering
Cryostat: 4 Al shields
Liquid Rb Scoh(Q,w)
Question for you:
How to remove this
contribution ?
Furnace: 12 shields (Al, Nb)
Let's simulate this...
Simulated total signal
Log scales

---
Sample environments: removing some background
Cryostat: 4 Al shields
Furnace: 12 shields (Al, Nb)
With ROC, mainly remains
container. Physics becomes
visible.

---
Neutron detectors
Detector model for MWPC: visualize (p,t) charge drift.
Can study:
detector spatial resolution
estimate detector saturation
background generated from detector housing.
Detection area 1x1 cm, He 5 bars, CF4 1 bar.
n
Resolution 0.4 cm

---
Virtual experiment: definition
Virtual experiment = simulation of sample + simulation of instrument
size 12{ { vec {F}}= - nabla u left ( { { size 8{r}}} right ) \( \) ital " {"} {}
nMoldyn
R, v
S(Q,w)
Pseudo-potential
This step is done prior to the virtual experiment.
Computationally intensive (e.g. Use clusters).
McStas provides a few sample S(Q,w).
Materials I'm working on: Rb, Au, Al, Bi, Hg, In, Pb, Se, Si, Ti, Tl
Molecular
Dynamics
Inter-atomic
potential u(r)
Ab-initio
(VASP)
Experiments
(corrected)

---
Virtual experiment: l-Rb
Experiments: Copley (PRA 1974) TOF and TAS
Simulation: 520 atoms in classical MD for 50 ps
T=315 K, c=1300 m/s, D=0.23 Å2/ps
Comparison with raw results (to avoid data analysis bias)
E0=4 meV
E0=33 meV
Structure: OK
Dynamics vs raw
experimental results:
OK
with container and
multiple scattering.
Parameters:
HybridTOF geometry
CP5 reactor power
Kahl 2-body potential
r=1.3 Å value for Rb

---
Example: liquid rubidium
COMPONENT Sample=Isotropic_Sqw(
  radius_o = 0.01, yheight = 0.055,
  Sqw_coh=”Rb_liq_coh.sqw”, Sqw_inc=”Rb_liq_inc.sqw”)
AT (0, 0, 0) RELATIVE Sample_pos
TRACE
COMPONENT Source
COMPONENT Sample=Isotropic_Sqw()
  AT 1.5 m RELATIVE Source
COMPONENT Collimator_radial
COMPONENT MonitorTotal WHEN SCATTERED
COMPONENT MonitorCoh WHEN coherent
COMPONENT MonitorInc WHEN incoherent
COMPONENT MonitorMulti WHEN SCATTERED>1
END
TOF E0=5 meV. l-Rb
McStas code (schematic)
We can separate contributions
from each part of the simulation
(not possible in real life)

---
Example: liquid indium
levitation
furnace+laser
f=2 mm
no container
simulation with
absolute flux
Indium: sabs=193 b ; sscatt=2.6 b
very few previous studies
Experiment performed on IN22@ILL TAS
a 'virtual vs real experiment'
Simulation: ab-initio using VASP, 200 atoms for 36 ps

---
Example: liquid indium
B4C nozzle
Xtal ?
Q [A-1]
IN22@ILL 1 week
Simulation: 1 day
Not yet perfect: simulations of B4C nozzle must be improved.
Virtual sample may be slightly crystallized.
Main features and absolute intensity reproduced.
Without simulation, the experiment would have been a failure.